## Slide 1: The use of HbA1c in the diagnosis of Type 2 Diabetes Mellitus: A National Audit of practice

## Slide 2: How did the audit come about?
Discussion around practice at the NCBAG
Sarah mentioned thinking about auditing it within her Trust………
! Let’s do a national audit
Sarah should do it
Sarah phoned a few friends……
85 responses to the survey were received

## Slide 3: Thankfully all respondents were aware of the WHO report from 2011

## Slide 4
WHO report summary
HbA1c can be used as a diagnostic test for diabetes provided that:
  Stringent quality assurance tests are in place
  Assays are standardised and aligned to international reference values
  No conditions are present that preclude its accurate measurement.
An HbA1c of ≥48mmol/mol (6.5%) is the cut point for diagnosing diabetes

## Slide 5: Is HbA1c is being used for diagnosis of Type 2 Diabetes Mellitus amongst users of your service?
Chart: 
Chart: 
| Yes | Think it is happening but not yet confirmed | No (Paediatrics) | No (others) | Answer not given
 | 79 | 11 | 4 | 2 | 5

## Slide 6: Do you promote the use of HbA1c for the diagnosis of Diabetes Mellitus?
Chart: 
| Neither promote nor discourage | Actively promote | Caveats | Paediatrics | Finances | In agreement with Diabetologists | No National Guidelines
Sales | 54 | 24 | 2 | 1 | 2 | 1 | 1

## Slide 7: Engagement with users about the use of HbA1c in diagnosis of Diabetes Mellitus
Chart: Number of responders/engagers
| Not Specified | Other | Discussion with Specialists | Trust Presentations | Handbook | Primary Care Presentations | Algorithms | Guidelines | Newsletter | Comments | Number of Engagers
 | 7 | 1 | 5 | 8 | 11 | 13 | 13 | 14 | 14 | 20 | 37

## Slide 8: Which groups have been involved in agreeing the content of presentations/publications?
8 respondents gave no answer
4 respondents only involved the laboratory staff
Nobody involved the patients!
Chart: Number of respondants (%)
Chart: Number of respondents
| Patient Groups | IT | CCGs | Individual GPs | Diabetologists | Laboratory Staff
 | 0 | 6 | 9 | 11 | 25 | 28

## Slide 9
Which of the following situations have you highlighted to users as inappropriate to use HbA1c for diagnosis?
Chart: 
| None of the above | Hypopituitarism | Other | Liver failure | HIV infection | Acute pancreatic damage or surgery | During acute illness | Treatment with corticosteroids or anti-psychotic therapy | Renal failure | Anaemia | Short duration of diabetic symptoms | All children and young people | Symptoms suggesting Type 1 DM | Pregnancy and up to 2 months post partum | Haemoglobinopathies (variants)
 | 0 | 7.1429 | 17.8571 | 21.4286 | 35.7143 | 60.7143 | 60.7143 | 60.7143 | 71.4286 | 82.1429 | 85.7143 | 89.2857 | 92.8571 | 96.4286 | 100
% respondents
Hypopituitarism was added
as a red herring to assess whether
respondents were merely ticking all
boxes.

## Slide 10: Identification of variants
Method | Number of laboratories | Does the method identify variants? |  | Does the method identify variant type? |  | 
 |  | Yes | No | Yes | No | Unknown
Affinity Primus | 3 | 2 |  | 1 | - | -
HPLC 8160 | 8 | 8 | 0 |  | 5 | 1
Electrophoresis – Capillarys 2 | 3 | 3 | 0 |  | 2 | -
HPLC D10 | 1 | 0 | 1 | - | - | -
HPLC G8 | 31 | 30 |  | 17 |  | 1
HPLC Variant II/Turbo | 10 | 9 |  |  | 6 | -
HPLC G7 | 7 | 6 |  | 5 |  | -
Affinity Hb-9210 | 9 |  | 8 |  |  | -
HPLC HA8180 | 5 | 4 |  | 2 |  | 1
Other | 4 | - | - | - | - | -
No answer | 10 | - | - | 1 | - | -
1
1
1
1
1
1
3
2
1
6
1
1
1
1
Identification of variants

## Slide 11: Dealing with variants
37 respondents would not report an HbA1c if a variant was detected
The following action is taken:
Chart: 
| None | Send a serum sample for fructosamine | Send sample for analysis of variant | Other | Suggest sample is sent for frustosamine | Send sample for analysis by another method | Suggest that a sample is sent for analysis of variant | Append a comment
 | 0 | 6 | 10 | 15 | 16 | 20 | 25 | 39
Number of responses

## Slide 12: Other actions for dealing with variants
Chart: 
| Suggest send sample for fasting glucose (if for diagnosis) | Speak to clinician | Add comment indicating HbA1c is unsuitable (if for diagnosis) | Suggest OGTT (if for diagnosis) | Review haematology/discuss with haematologist | Depends on variant and interference | Depends on purpose of HbA1c
Number | 1 | 1 | 1 | 3 | 3 | 7 | 10
Number of responses

## Slide 13: Question 17 asked for running CVs for glucose and HbA1c methods
Guidelines and Recommendations for Laboratory Analysis in the Diagnosis and Management of Diabetes Mellitus. David B. Sacks, David E. Bruns, David E. Goldstein, Noel K. Maclaren, Jay M. McDonald and Marian Parrott Clinical Chemistry 48:3 436–472 (2002)
The intra-laboratory CV for plasma glucose should be < 2.2%
The intra-laboratory CV for glycated HbA1c should be < 2%

## Slide 14
Running CVs for glucose and HbA1c IQC
HbA1c
Glucose

## Slide 15
Chart: Glucose vs CV Data

## Slide 16
Chart: HbA1c vs CV Data

## Slide 17: Do you have a mechanism to distinguish between HbA1c for diagnosis and HbA1c for monitoring of diabetes?
65% of respondents are not able to distinguish
Those that could used the following:
Mechanism | Number
Clinical details | 13
Known Type 2 Diabetes patients flagged in system | 2
Manually distinguished at reporting | 11
Separate request codes | 10
Other | 4

## Slide 18: Dealing with HbA1c Results
Chart: 
| Validation | Reference Range | Phone limits
Yes | 39 | 42 | 15
No | 32 | 24 | 52
No Response | 13 | 18 | 17

## Slide 19: HbA1c reference ranges
56% quote 20-41 mmol/mol
15% quote <42 mmol/mol
7% quote <48 mmol/mol
The remainder all have their own view!
  e.g. <53, 27-47, 30-42 mmol/mol

## Slide 20: Dealing with HbA1c Results
Chart: 
| Validation | Reference Range | Phone limits
Yes | 39 | 42 | 15
No | 32 | 24 | 52
No Response | 13 | 18 | 17

## Slide 21
Interpretative comments
≥48 mmol/mol:
Mostly mention diagnostic cut-off, need for symptoms or repeat sample
21% do not mention diagnostic use of HbA1c
42-47 mmol/mol:
58% indicate patient has increased risk of developing diabetes
<42 mmol/mol:
57% state result is not consistent with diabetes, most also specify patient may be at risk of developing diabetes
Several labs use a standard comment for all HbA1c results regardless of value

## Slide 22: Demand Management
30% have a demand management system for HbA1c in place
This is either achieved at requesting, booking in or requests being referred to a manual vetting queue
Response | No. of responses | Notes
28 days | 1 | 
30 days/1 month | 2 | 
56 days | 4 | 1 will reinstate if sample is diagnostic
6 weeks | 2 | 
60 days/2 months | 8 | 3 check whether sample is diagnostic
72 days | 1 | 

## Slide 23: Do you or your users continue to carry out glucose tolerance tests?
Chart: 
| Yes, in certain patient groups | Yes, GTTs and HbA1cs are being requested in the same patients | No | No answer
 | 49 | 17 | 2 | 17

## Slide 24: If GTTs are only being carried out in certain patient groups, please indicate which ones
Chart: 
| Other, please specify | Patients with renal failure | Pateints with liver failure | Not applicable | Patients with anaemia | Patients with known haemoglobinopathies (variants) | Investigation of gestational diabetes
 | 7 | 3 | 4 | 4 | 5 | 13 | 45

## Slide 25: GTTs in other patient groups
Chart: 
| Other | Confirm diagnosis | HbA1c not appropriate | GP/out-patients
Series 1 | 2 | 5 | 6 | 14

## Slide 26
Chart: No. requests per 100,000
HbA1c 2010 | 2911.3636 | 4278.2222 | 4757.25 | 5455.6 | 5666.6667 | 7197.1795 | 7326.3077 | 7484 | 7599.6 | 7759.6667 | 7777.7778 | 7919.3333 | 7950 | 8135 | 8307.0961 | 8716.4706 | 9415.4472 | 9448.8571 | 9511.6129 | 9687.5 | 9813.6364 | 10103.6111 | 10856.087 | 11038.1902 | 11099.6923 | 11106.7568 | 11200 | 11226.2554 | 12152.6667 | 12451 | 13027.1698 | 13072.407 | 13237.4286 | 13700.4 | 19319.8113 | 19375.8621 | 20241.6 | 21576.4
HbA1c 2013 | 20658.9773 | 7534 | 6325.25 | 7223.2 | 15000 | 10305.1282 | 13944.7692 | 17112 | 12494.4 | 10810 | 10000 | 11586.4444 | 11556.2667 | 13288 | 21015.329 | 12065.5882 | 11748.7805 | 12309.7143 | 13341.0753 | 31562.5 | 13910.9091 | 10913.6111 | 12585.6522 | 14505.8282 | 15283.0769 | 25800.5405 | 19842.1053 | 18883.2481 | 24613.5556 | 14974 | 19511.8868 | 12214.0495 | 17154.8571 | 19771.6 | 27785.8491 | 35772.7586 | 27078.4 | 18306.8
Total number of HbA1c analyses per 100,000 GP population

## Slide 27: Total Number of GTT requests per 100,000 GP population
Chart: No. requests per 100,000
| Princess Royal Hospital | Royal Victoria Hospital | Southmead Hospital | Croydon University Hospital | Aintree University Hospital | Hope Hospital | The York Hospital | Ulster Hospital | Royal Blackburn Hospital | Victoria Hospital | Leeds General Infirmary | Torbay Hospital | The Princess Alexandra Hospital NHS Trust | Nevill Hall Hospital | Raigmore Hospital | Arrowe Park Hospital | Huddersfield Royal Infirmary | Alexandra Hospital | Queen's Hospital | Queen's Hospital | Gloucestershire Royal Hospital | Royal United Hospital | Cumberland Infirmary | Sandwell and West Birmingham Hospitals NHS Trust | Royal Preston Hospital | Royal Liverpool University Hospital | Kent & Canterbury Hospital | Sherwood Forest Hospitals NHS Foundation Trust | Darlington Memorial Hospital | 26 | 19 | 18 | 11 | 38 | 22 | 13 | 32 | 05 | 36 | 29 | 34 | 02 | 17 | 08 | 14 | 09 | 10 | 27 | 06 | 30 | 28 | 15 | 21 | 04 | 01 | 25 | 20 | 37
GTT 2010 | 161.2 | 248.8889 | 324.1667 | 385.6369 | 390.6061 | 402 | 413.5556 | 424 | 541.5094 | 554.7222 | 587.3864 | 604.5614 | 611.25 | 780.1471 | 810 | 824.2857 | 899.2308 | 943.8512 | 1018.8 | 1037.5714 | 1057.2 | 1064.4444 | 1093.75 | 1430.7547 | 1462.4324 | 1692.8 | 1996.4615 | 2006.2419 | 3550
GTT 2013 | 469.6 | 374.2222 | 445.6667 | 483.7398 | 308.4848 | 178.4 | 367.5556 | 917 | 943.3962 | 619.4444 | 485.7955 | 318.9474 | 481.25 | 422.5 | 698.6957 | 740.5714 | 808.3077 | 664.8295 | 966 | 1020.4286 | 577.4667 | 794.2222 | 437.5 | 1018.6792 | 641.6216 | 1112 | 1932.4615 | 1534.129 | 1249.8333

## Slide 28
Key Findings
HbA1c is widely used for diagnosis, but this is not being guided
Engagement with users, including patients, could be improved
Caveats of using HbA1c are highlighted but inconsistent
Not all laboratories are aware of their method characteristics
CVs for HbA1c are not meeting recommendations
There is variability around commenting
There remains a mixture of glucose and HbA1c measurements, which aren’t always under the control of the laboratory

## Slide 29
Recommendations
Follow the WHO guidance
Engage with users
Highlight caveats
Know your method
Review comments
Highlight differences between GTT and HbA1c

## Slide 30
Acknowledgements
National Clinical Biochemistry Audit Group
Team North West:
Shirley Bowles
Sally Hanton
Andrew Hutchesson
Mandy Pickersgill
Tony Tetlow
Advisors:
Garry John and Emma English
David Oleesky
Adrian Heald
Tom Hutchinson
All survey respondents